OPENING Cs SOURCE CHECKLIST

___	Set HV, Ionizer, & Reservoir to 0 in Sources window

___	Flip off Cs power supply switch in back of cabinet

___	Disconnect following 3 feedthroughs on Cs chassis:

___	Ionizer/reservoir

___	Cs HV

___	LDuo

___	Disconnect following 3 pieces on Wien filter chassis:

___	Wien filter itself feedthrough

___	Some random deflector (CDuo?) feedthrough

___	Water line clamp

___	Disconnect water line from Cs source (push in to disengage, then pull out)

___	Reconnect water line loop w/o Cs source so pumps get water

___	Vent source chamber: either in Vacuum window or from cabinet

___	Take off Wien filter (using gloves from now on)

___	Remove screws on Wien filter chassis

___	Carefully lift and remove Wien filter & place it aside (don’t drop!)

___	Looking through Wien filter hole & check to see Cs source can clear interior wall

___	If not, use Cs adjustment dial to move into a safe position before removal

___	Take off Cs source

___	Remove all screws, careful not to let source fall out

___	Grasping Cs source by black handle, gracefully remove Cs source

___	Place it on bench for detailed look