Cs SOURCE STARTUP CHECKLIST

HARDWARE

___    Is water line (black w/ gray connectors) on back of instrument hooked up?

___    Is water fan (the top one) in closet spinning?
          ___    If bottom is spinning too, unplug white plug (freon line)

___    Are the following vacuum/safety lights on cabinet green ?
          ___    Big electronics button above the key (if not, turn key to "on")
          ___    Multicollection
          ___    Central
          ___    Vacuum safety (lowest one, there are two)
          ___    Hv pol pos (if not, proceed anyway)

___    Is the fault light off? (if not, there is a vacuum error: fix it)


SOFTWARE

___    Is HV on?

___    Are all the usual buttons green?
          ___    Setup
          ___    Preset
          ___    Sources
          ___    Tuning
          ___    NMR
          ___    Holder
          ___    CCD
         
___    Load most recent Cs .isf file in 'Presets' and apply them

___    In 'Sources', is polarity positive and 'Cs' selected? (If not, select them)

___    Automatic
          ___    Press play and enter appropriate resevoir, ionizer, and HV values
          ___    Come back in 30 minutes

___    Manual (Do the following in order)
          ___    Increase the HV in steps up to 8kV
          ___    Increase ionizer to desired value gradually
          ___    Wait for at least 10 minutes to heat
          ___    Increase resevoir to desired value gradually
___    Are resevoir and ionizer lights green? (if not, something is wrong...likely pressure)